## Slide 1
104
150pF
10nF
104
Sens_Out
V+(5V)
100Ω
100Ω
GND
PGA309
Sensor device
Top side
Bottom seide
Via
PGA309 Sensor board (without EEPROM)
4
Create the evaluation board
Install PGA309
Install sensor (P3->spec)
Force 5V to V+ and GND
Measure between Vexc and GND voltage, it was 0V.
Measure between Vsd and GND, Vsa and GND, there were 5V.
Measure between Out and GND was 0V.
Why is the voltage is not supplied to a sensor?
Unless it writes something in EEPROM, doesn't PGA309 work?

## Slide 2
104
150pF
10nF
104
Sens_Out
V+(5V)
100Ω
100Ω
GND
PGA309
Sensor device
PGA309 Sensor Board (with EEPROM)
4
Why is the voltage is not supplied to a sensor?
Unless it writes something in EEPROM, doesn't PGA309 work?
Would you please teach to me the method of EEPROM programing and programing data.
Do you have any document?
EEPROM
Top side
Bottom seide
Via
Create the evaluation board
Install PGA309
Install sensor (P3->spec)
Install blank EEPROM
Force 5V to V+ and GND
Measure between Vexc and GND voltage, it was 0V.
Measure between Vsd and GND, Vsa and GND, there were 5V.
Measure between Out and GND was 0V.

## Slide 3
I would like to do the following function by using PGA309.
4
Sensor output
センサ素子出力の近似式
I would like to do the following function;
/ Linearization, Amplification, Offset compensation
/ Temperature drift compensation of output offset, Correction for temperature of sensitivity
Question:
/ Please let you know to me a flow to make PGA309 work.
/ When using PGA309 EVM-USB, which button and tab should be operated and
order?
/ When working EVM using our company's sensor, where should we just input for linearization and correction of temperature?
Please teach me concretely.
Differential output between Vin1 and Vin2 input(V)
Load (full scale)
Load
Calculation of the approximate value of a sensor output