2.ソフトウェア使用方法

How to use the software

ソフトウェア 「 PGA305EVM 」 を起動する

Start the software "PGA305EVM"

1

2.ソフトウェア使用方法

How to use the software

1.Interface Setting

起動後 「 ◆Interface Settings 」 の画面（下図）が表示される

画面右上 「 Interface Disconnected? 」 のチェックがついていればクリックしてチェックを外す画面左上の 「 I2C 」 をクリック

If "Interface Disconnected?" Is checked at the top right of the screen, click it to uncheck it.

Click "I2C" at the top left of the screen

2.ソフトウェア使用方法

1.Interface Setting

画面左上 「 Compensation Mode 」 → 「 Digital Interface 」画面右上 「 DI Not Active 」 → 「 DI Active 」

画面右下 「 HW DISCONNECTED 」 → 「 HW CONNECTED 」 Confirm that it has changed and proceed to the next.

変更しない

2. How to use the software

2.EEPROM Config-Analog

画面左 from「 ◇Device Calibration 」 select to「 ◆EEPROM Config-Analog 」

4

2. How to use the software

2.EEPROM Config-Analog

画面中央 「 EEPROM Register Details 」 の 「 Write Data 」 Click the field and select each item

項目	選択値

DAC_RATIO	Absolute Mode

DAC_GAIN	4.00 V

4_20MA_EN	No

DACCAP_EN	Yes

VBRDG_CTRL	2.5 V

P_GAIN	40.00

P_INV	No

T_GAIN	1.33

T_INV	No

TEMP_MUX_CTRL	VINTP-VINTN

ITEMP_CTRL	OFF

TSEM_N	Differential

5

2. How to use the software

2.EEPROM Config-Analog

After selecting each item, check the "Register Name" field and make all items selected.

With all items selected, click "Write Selected" at the bottom center of the screen to write.

6

2. How to use the software

2.EEPROM Config-Analog

After writing, turn off the power of the product → Power on

After turning on the power, click "I2C" to communicate with the connected water level gauge.

Check the "Register Name" field and click "Read Selected"

7

2. How to use the software

2.EEPROM Config-Analog

After confirming that "Write Data" and "Read Data" are the same as shown in the figure below,

proceed to the next step.

8

2. How to use the software

3.ADC & DAC Calibration

From "◇ Device Calibration" on the left side of the screen,

Select "◆ ADC & DAC Calibration"

9

2. How to use the software

3.ADC & DAC Calibration

Enter "DAC Codes" in the column from top to bottom　"51E" "11EB" "1EB8"「 DAC Codes 」

Enter "DAC Codes" in the column from top to bottom,Enter "0.4000", "1.4000", "2.4000"

51E 11EB 1EB8

0.4000 1.4000 2.4000

10

2. How to use the software

3.ADC & DAC Calibration

Click "DAC Code" under "DAC Code Calculation"

Click "Write DAC Code" when the frame color changes to blue.

0.4000

1.4000 2.4000

11

2. How to use the software

3.ADC & DAC Calibration

Check the value displayed on the multimeter measuring the output voltage of the product.

Operate in the same way for each column (51E, 11EB, 1EB8) of "DAC Code" and check the value.

mV

0.4132 V

12