## Slide 1
Calibration results

## Slide 2
Temperature calibration results
1. Calibration condition
- Program Version : Multi-Cal 2.3(Beta)
- Temperature: 3-point(-10℃, 25℃, 80℃)
- Calibration model/QTY : Gauge Pressure 10bar, Iout type / 64EA
- Model File set up : Temp ADC Resolution 15bit, Vs 5V
- DMM Scripts : Pause 1.5(recommended by Ian)

## Slide 3
Temperature calibration results
2. Temperature calibration results
- Temperature calibration(fail/Total) : 5EA/64EA
- Save calibration data on EEPROM of device: We found all failure data(64EA) on EEPROM when checking on EEPROM Data Read
- Save calibration data on PC(.CSV) : Normal
- Convert calibration data to text file : Normal
Complete saving on PC
Complete converting to Text file

## Slide 4
Problems during calibration

## Slide 5
Problem 1. Temperature calibration procedure failure
1. Temperature calibration procedure failure – CANNOT DUPLICATE ISSUE
- Set program to proceed from -10℃(First), 25℃(Second) to 80℃(Third)
- But, after -10℃ calibration, move to 80℃(But temperature setting of equipment is 25℃)
- Second(actually 25℃) calibration data is written on 80℃ table of program.
- Check that data on ‘Find Vin_Low’ is recorded on 80℃ table.(For normal case, it should be recorded on 25 ℃ table)
- Stop calibration and click ‘Re-Cal Selected Temperature’ to change temperature to 25℃ and start calibration again.
- After High/Low calibration, program complete calibration without third calibration(80℃ calibration) (Save data on EEPROM and PC)
- Currently, we thinks this is program bug.

## Slide 6
Problem 2. EEPROM data saving failure
2. EEPROM data saving failure
- Check data on EEPROM after completing calibration.
- When reading ch.0(complete temperature calibration), we found ‘OAOA’.
- All data is OAOA.
Select/Read Ch.0
All data is ‘OAOA’

## Slide 7
Problem 3. Data Read failure
3. try to write PC data on EEPROM and read it. Still can find failure data
- Complete saving data on PC(.CSV) and converting Text file
- Select Ch.0 and rewrite PC data(It was ok) on EEPROM
Select/write Ch.0 data

## Slide 8
Problem 3. Data Read failure
- Complete Writing normally.
same
Ch.0 data(text file) of PC
Data on EEPROM

## Slide 9
Problem 3. Data Read failure
- After completing to Write, read ch.0 EEPROM again.
- Even if writing PC data on EEPROM, all data is ‘OAOA’ after reading ch.0
Different
Ch.0 data(text file) of PC
Data on EEPROM

## Slide 10
Conclusion

## Slide 11
Conclusion
1. Use Multi-Cal Version 2.1 again
- It seems like that there is severe bug on Multi-Cal Version 2.3(beta) as a result of 64EA 3-point temperature calibration.
- Currently Autonics got over 2Ku order so start temperature calibration immediately.
- So, determine to use 2.1 Version until revising Multi-Cal 2.3 for calibration
- Need your strong support to release official Version 2.3 soon.