I get the same results if I am using the one wire on the programing pin tied to vout or just four wires with programing pin ties high. See wiring below:

Today I changed C3 to 220pF and then I removed C3 completely, I got the same results. I ordered a 100pF and it will be here in the morning. I will install and test with it in the AM.

I have also attached a pre-cal file/ post cal file

Also, I measured the raw sensor output (VEXE seems to work, and I get ~2mV to 8.5mV FS). I can change the FS setting and the calibration follows i.e.  0.5 to 2V out I get 1.97 to 2V out if I want 0.5 to 1 V out I get .97 to 1 V out.  The odd thing is that it does respond to pressure and the communications seem fine.