USB-DAQ EVM  offset images.

The above image shows that the target set point is 4.5v. The program believes that it achieved an output of 4.97v. The actual output is 4.584v.

The above image shows a target set point of .5v.  the program shows that it achieved .50395v when the actual voltage is .518v

With a Target voltage offset to 4.407v the program believes it achieved 4.4046v. The actual output is 4.49v

Here the Target voltage is .484v and the program thinks it achieved .48773v. The actual output is .501v