## Sheet1 (English)
Digital Power meter （WT110）
| Voltage(V) | Current(A) | Power(W) | Frequency(Hz) | Power Factor(PF) | Reactive power(var) | Apparent power(VA)
Display value | 100 | 1.043 | 64.8 | 50 | 0.621 | 81.9 | 104.5 | Active power：
Theoretical value | 100 | 1.043 | 64.7703 | 50 | 0.6201 | 81.7514 | 104.4349
Reactive power：
TI-EVM GUI (EVM430-F6779)
| Voltage(V) | Current(A) | Power(W) | Frequency(Hz) | Power Factor(PF) | Reactive power(var) | Apparent power(VA) | Apparent power：
Display value | 99.985 | 1.046 | 63.8192 | 50 | 0.8583 | 38.1394 | 74.3472
Theoretical value | 99.985 | 1.046 | 89.7647 | 50 | 0.85839 | 53.6673 | 74.3472
Theoretical value2 | 99.985 | 1.046 | 63.8192 | 50 | 0.61022 | 82.8552 | 104.5843 | From the Pythagorean theorem
Apparent power^2 = Active power^2 + Reactive power^2 から
About theoretical formulas | Active power = voltage x current x power factor (Cosθ) | [6 empty] | Apparent power=root(（Active power^2)+(Reactive power^2)
Reactive power = voltage x current x inefficiency (Sinθ) = voltage x current x (root (1-power factor^2))
Apparent power = root((active power^2)+(reactive power^2))
About theoretical value2 | Since the active power has the same value as the power meter, calculate the power factor from the active power.
| Power meter | 64.8
| TI-EVM GUI | 63.8192
From, active power = voltage x current x power factor (Cosθ),
Power factor (Cosθ) = active power/(voltage x current)
| Recalculated value | 0.61022 |  | Power meter | 0.621
|  |  |  | TI-EVM GUI | 0.8583 | *There are differences from the theoretical formula results.
＜About calibration＞
There are three setting items to calibrate: voltage, current, and phase.
Only GAIN and Phase are subject to EVM (meter side) due to calibration.
Verify that the power factor value does not change if the calibration phase is adjusted to the wrong value.
However, we confirmed that the values of active power, reactive power, and apparent power vary depending on the setting values.
*We understand that the correct way to do this is to change the phase using equipment used by power meter certification bodies.
However, please note that we are testing the movement by shifting the phase of the GUI side in a temporary environment.
Perform calibration