## Sheet1
2020.9.24
Research & Development Dept.Ⅱ J.MORITA MFG. CORP.
input current waveform at a load connected.(20200924)
Regarding UCC28019A, when the input is 200V or more, the input current waveform is distorted at the load connected.
Waveforms were acquired with adding a bypass diode(parallel with L2-D6)
(Could you please see the schematic below)
■Conclsion
| ・ | Adding a bypass diode does not improve the current distortion when the load is conncted.
→ Fig1 shows the inrush current waveform when the power is turned on near the input voltage peak and the waveform with the load connected about 2.5 seconds after that.
Since the time from turning on the power to turning on the load is several seconds or more, it is considered that the inductor will not be saturated due to the inrush current.
(At this time, I connected the load about 2.5 seconds after the power was turned on. It can't be faster than this.)
→ Fig2 shows the distorted input current waveform when the load is connected and the expanded waveform.
It was assumed that the inductor would be saturated due to the effect of the inrush current when the power was turned on, but there was no sign of saturation when looking at the shunt current waveform at the switching frequency.
Circuit diagram
■Input: 230V/50Hz
Fig.1
Waveform when the power is turned on and the load is turned on. (Input Voltage, Bypass diode current, Input current)
Adding a bypass diode does not improve the current distortion when the load is applied.
It takes at least about 2 seconds from turning on the power to turning on the load, so it is considered that there is no effect on the saturation of the inductor.
Fig.2
Waveform when the load is turned on. (Input Voltage, Shunt current, Input current)
When the load was connected, the inductor current waveform was not saturated even if the shunt current waveform was expanded.
UCC28019ADG4
Power ON
100k
27mΩ
4700p
220
1.1M
1.8M
1.8M
1.8M
未実装
0.68u
1000p
220p
1u
0.1u
820p
0.22u
2.2u
1u
33.2k
1.5u
35V 100u
未実装
未実装
AC100～264V
Input current
Bypass diode
current
10A
-10A
500V
Input voltage
Input Current
Shunt current
Input voltage
ZOOM
ZOOM
bypass-diode current
added bypass-diode
Load ON
ZOOM
Input current
Shunt current
Input voltage
-500V
10A
Load ON
ZOOM
The inductance waveform is not satulated.
10us
500V
-500V
10A
-10A
10A
2.5sec