## input
* Texas Instruments Data Flash File
* File created Wed Apr 02 10:45:51 2025
*
* Device Number 100
* Firmware Version 0.16
* Build Number not available
* Order Number not available
*
* bqz Device Number 100
* bqz Firmware Version 0.16
* bqz Build Number 17
*
* Field Order: Class name | Subclass name | Parameter name | Parameter Value | Display Units
Configuration | Safety | OT Chg | 55 | 1degC
Configuration | Safety | OT Chg Time | 2 | Seconds
Configuration | Safety | OT Chg Recovery | 50 | 1degC
Configuration | Safety | OT Dsg | 60 | 1degC
Configuration | Safety | OT Dsg Time | 2 | Seconds
Configuration | Safety | OT Dsg Recovery | 55 | 1degC
Configuration | Charge Inhibit Cfg | Chg Inhibit Temp Low | 0 | 1degC
Configuration | Charge Inhibit Cfg | Chg Inhibit Temp High | 45 | 1degC
Configuration | Charge Inhibit Cfg | Temp Hys | 5 | 1degC
Configuration | Charge | Suspend Low Temp | -5 | 1degC
Configuration | Charge | Suspend High Temp | 55 | 1degC
Configuration | Charge | Pb Temp Comp | 24.96 | %
Configuration | Charge | Pb Reduction Rate | 10 | %
Configuration | Charge Termination | Taper Current | 100 | mAmp
Configuration | Charge Termination | Min Taper Capacity | 25 | mAmpHr
Configuration | Charge Termination | Cell Taper Voltage | 200 | mVolt
Configuration | Charge Termination | Current Taper Window | 40 | Seconds
Configuration | Charge Termination | TCA Set % | 99 | Percent
Configuration | Charge Termination | TCA Clear % | 95 | Percent
Configuration | Charge Termination | FC Set % | 100 | Percent
Configuration | Charge Termination | FC Clear % | 98 | Percent
Configuration | Charge Termination | DODatEOC Delta T | 10 | 1degC
Configuration | Charge Termination | NiMH Delta Temp | 3 | 1degC
Configuration | Charge Termination | NiMH Delta Temp Time | 180 | Seconds
Configuration | Charge Termination | NiMH Hold Off Time | 100 | Seconds
Configuration | Charge Termination | NiMH Hold Off Current | 240 | mAmp
Configuration | Charge Termination | NiMH Hold Off Temp | 25 | 1degC
Configuration | Charge Termination | NiMH Cell Negative Delta Volt | 17 | mVolt
Configuration | Charge Termination | NiMH Cell Negative Delta Time | 16 | Seconds
Configuration | Charge Termination | NiMH Cell Neg Delta Qual Volt | 4200 | mVolt
Configuration | Data | Manuf Date | 1980-1-1 | Day + Mo*32 + (Yr -1980)*256
Configuration | Data | Ser. Num. | 1 | hex
Configuration | Data | Cycle Count | 3 | Count
Configuration | Data | CC Threshold | 8100 | mAmpHr
Configuration | Data | Max Error Limit | 20 | %
Configuration | Data | Design Capacity | 9000 | MilliAmpHour
Configuration | Data | Design Energy | 32400 | MilliWattHour
Configuration | Data | SOH Load I | -400 | MilliAmp
Configuration | Data | Cell Charge Voltage T1-T2 | 4200 | mV
Configuration | Data | Cell Charge Voltage T2-T3 | 4200 | mV
Configuration | Data | Cell Charge Voltage T3-T4 | 4100 | mV
Configuration | Data | Charge Current T1-T2 | 10 | Percent
Configuration | Data | Charge Current T2-T3 | 50 | Percent
Configuration | Data | Charge Current T3-T4 | 30 | Percent
Configuration | Data | JEITA T1 | -10 | degC
Configuration | Data | JEITA T2 | 10 | degC
Configuration | Data | JEITA T3 | 45 | degC
Configuration | Data | JEITA T4 | 55 | degC
Configuration | Data | Design Energy Scale | 1 | Number
Configuration | Data | Device Name | bq34z100-G1 | -
Configuration | Data | Manufacturer Name | Texas Inst. | -
Configuration | Data | Device Chemistry | LION | -
Configuration | Discharge | SOC1 Set Threshold | 150 | mAh
Configuration | Discharge | SOC1 Clear Threshold | 175 | mAh
Configuration | Discharge | SOCF Set Threshold | 75 | mAh
Configuration | Discharge | SOCF Clear Threshold | 100 | mAh
Configuration | Discharge | Cell BL Set Volt Threshold | 0 | mVolt
Configuration | Discharge | Cell BL Set Volt Time | 0 | Seconds
Configuration | Discharge | Cell BL Clear Volt Threshold | 5 | mVolt
Configuration | Discharge | Cell BH Set Volt Threshold | 4300 | mVolt
Configuration | Discharge | Cell BH Volt Time | 2 | Seconds
Configuration | Discharge | Cell BH Clear Volt Threshold | 5 | mVolt
Configuration | Discharge | Cycle Delta | 0.05 | %
Configuration | Manufacturer Data | Pack Lot Code | 0 | hex
Configuration | Manufacturer Data | PCB Lot Code | 0 | hex
Configuration | Manufacturer Data | Firmware Version | 0 | hex
Configuration | Manufacturer Data | Hardware Revision | 0 | hex
Configuration | Manufacturer Data | Cell Revision | 0 | hex
Configuration | Manufacturer Data | DF Config Version | 0 | hex
Configuration | Integrity Data | Static Chem DF Checksum | 0 | Number
Configuration | Lifetime Data | Lifetime Max Temp | 32.9 | 1degC
Configuration | Lifetime Data | Lifetime Min Temp | 20 | 1degC
Configuration | Lifetime Data | Lifetime Max Chg Current | 1810 | mAmp
Configuration | Lifetime Data | Lifetime Max Dsg Current | -1828 | mA
Configuration | Lifetime Data | Lifetime Max Pack Voltage | 449 | 20mV
Configuration | Lifetime Data | Lifetime Min Pack Voltage | 252 | 20mV
Configuration | Lifetime Temp Samples | LT Flash Cnt | 50 | Count
Configuration | Registers | Pack Configuration | 161 | flags
Configuration | Registers | Pack Configuration B | ff | flags
Configuration | Registers | Pack Configuration C | 30 | flags
Configuration | Registers | LED_Comm Configuration | 0 | flags
Configuration | Registers | Alert Configuration | 0 | flags
Configuration | Registers | Number of series cell | 2 | num
Configuration | Lifetime Resolution | LT Temp Res | 1 | 1degC
Configuration | Lifetime Resolution | LT Cur Res | 100 | mA
Configuration | Lifetime Resolution | LT V Res | 1 | 20mV
Configuration | Lifetime Resolution | LT Update Time | 60 | Seconds
Configuration | LED Display | LED Hold Time | 4 | Num
Configuration | Power | Flash Update OK Cell Volt | 2800 | mVolt
Configuration | Power | Sleep Current | 10 | mAmp
Configuration | Power | FS Wait | 0 | Seconds
System Data | Manufacturer Info | Block A 0 | 0 | hex
System Data | Manufacturer Info | Block A 1 | 0 | hex
System Data | Manufacturer Info | Block A 2 | 0 | hex
System Data | Manufacturer Info | Block A 3 | 0 | hex
System Data | Manufacturer Info | Block A 4 | 0 | hex
System Data | Manufacturer Info | Block A 5 | 0 | hex
System Data | Manufacturer Info | Block A 6 | 0 | hex
System Data | Manufacturer Info | Block A 7 | 0 | hex
System Data | Manufacturer Info | Block A 8 | 0 | hex
System Data | Manufacturer Info | Block A 9 | 0 | hex
System Data | Manufacturer Info | Block A 10 | 0 | hex
System Data | Manufacturer Info | Block A 11 | 0 | hex
System Data | Manufacturer Info | Block A 12 | 0 | hex
System Data | Manufacturer Info | Block A 13 | 0 | hex
System Data | Manufacturer Info | Block A 14 | 0 | hex
System Data | Manufacturer Info | Block A 15 | 0 | hex
System Data | Manufacturer Info | Block A 16 | 0 | hex
System Data | Manufacturer Info | Block A 17 | 0 | hex
System Data | Manufacturer Info | Block A 18 | 0 | hex
System Data | Manufacturer Info | Block A 19 | 0 | hex
System Data | Manufacturer Info | Block A 20 | 0 | hex
System Data | Manufacturer Info | Block A 21 | 0 | hex
System Data | Manufacturer Info | Block A 22 | 0 | hex
System Data | Manufacturer Info | Block A 23 | 0 | hex
System Data | Manufacturer Info | Block A 24 | 0 | hex
System Data | Manufacturer Info | Block A 25 | 0 | hex
System Data | Manufacturer Info | Block A 26 | 0 | hex
System Data | Manufacturer Info | Block A 27 | 0 | hex
System Data | Manufacturer Info | Block A 28 | 0 | hex
System Data | Manufacturer Info | Block A 29 | 0 | hex
System Data | Manufacturer Info | Block A 30 | 0 | hex
System Data | Manufacturer Info | Block A 31 | 0 | hex
Gas Gauging | IT Cfg | Load Select | 1 | Number
Gas Gauging | IT Cfg | Load Mode | 0 | Number
Gas Gauging | IT Cfg | Max Res Factor | 50 | num
Gas Gauging | IT Cfg | Min Res Factor | 1 | num
Gas Gauging | IT Cfg | Ra Filter | 500 | num
Gas Gauging | IT Cfg | Min PassedChg NiMH-LA 1st Qmax | 50 | %
Gas Gauging | IT Cfg | Maximum Qmax Change | 100 | %
Gas Gauging | IT Cfg | Cell Terminate Voltage | 3100 | mVolt
Gas Gauging | IT Cfg | Cell Term V Delta | 200 | mVolt
Gas Gauging | IT Cfg | ResRelax Time | 500 | Seconds
Gas Gauging | IT Cfg | User Rate-mA | 0 | MilliAmp
Gas Gauging | IT Cfg | User Rate-Pwr | 0 | mW/cW
Gas Gauging | IT Cfg | Reserve Cap-mAh | 0 | MilliAmpHour
Gas Gauging | IT Cfg | Reserve Energy | 0 | mWh/cWh
Gas Gauging | IT Cfg | Max Scale Back Grid | 4 | num
Gas Gauging | IT Cfg | Cell Min DeltaV | 0 | mVolt
Gas Gauging | IT Cfg | Ra Max Delta | 15 | %
Gas Gauging | IT Cfg | Design Resistance | 42 | mOhms
Gas Gauging | IT Cfg | Reference Grid | 4 | -
Gas Gauging | IT Cfg | Qmax Max Delta % | 10 | mAmpHour
Gas Gauging | IT Cfg | Max Res Scale | 32000 | Num
Gas Gauging | IT Cfg | Min Res Scale | 1 | Num
Gas Gauging | IT Cfg | Fast Scale Start SOC | 10 | %
Gas Gauging | IT Cfg | Charge Hys V Shift | 40 | mVolt
Gas Gauging | IT Cfg | Smooth Relax Time | 1000 | s
Gas Gauging | Current Thresholds | Dsg Current Threshold | 300 | mAmp
Gas Gauging | Current Thresholds | Chg Current Threshold | 1000 | mAmp
Gas Gauging | Current Thresholds | Quit Current | 40 | mAmp
Gas Gauging | Current Thresholds | Dsg Relax Time | 300 | Seconds
Gas Gauging | Current Thresholds | Chg Relax Time | 60 | Seconds
Gas Gauging | Current Thresholds | Cell Max IR Correct | 400 | mV
Gas Gauging | State | Qmax Cell 0 | 9445 | mAmpHr
Gas Gauging | State | Cycle Count | 2 | num
Gas Gauging | State | Update Status | 5 | num
Gas Gauging | State | Cell V at Chg Term | 4167 | mVolt
Gas Gauging | State | Avg I Last Run | -299 | mAmp
Gas Gauging | State | Avg P Last Run | -1131 | MilliWattHour
Gas Gauging | State | Cell Delta Voltage | 2 | mVolt
Gas Gauging | State | T Rise | 20 | Num
Gas Gauging | State | T Time Constant | 1000 | Num
Ra Table | R_a0 | R_a0 Flag | ff55 | Hex
Ra Table | R_a0 | R_a0 0 | 88 | Num
Ra Table | R_a0 | R_a0 1 | 90 | Num
Ra Table | R_a0 | R_a0 2 | 102 | Num
Ra Table | R_a0 | R_a0 3 | 134 | Num
Ra Table | R_a0 | R_a0 4 | 119 | Num
Ra Table | R_a0 | R_a0 5 | 91 | Num
Ra Table | R_a0 | R_a0 6 | 106 | Num
Ra Table | R_a0 | R_a0 7 | 112 | Num
Ra Table | R_a0 | R_a0 8 | 107 | Num
Ra Table | R_a0 | R_a0 9 | 121 | Num
Ra Table | R_a0 | R_a0 10 | 147 | Num
Ra Table | R_a0 | R_a0 11 | 186 | Num
Ra Table | R_a0 | R_a0 12 | 269 | Num
Ra Table | R_a0 | R_a0 13 | 748 | Num
Ra Table | R_a0 | R_a0 14 | 7558 | Num
Ra Table | R_a0x | R_a0x Flag | ffff | Hex
Ra Table | R_a0x | R_a0x 0 | 88 | Num
Ra Table | R_a0x | R_a0x 1 | 90 | Num
Ra Table | R_a0x | R_a0x 2 | 102 | Num
Ra Table | R_a0x | R_a0x 3 | 134 | Num
Ra Table | R_a0x | R_a0x 4 | 119 | Num
Ra Table | R_a0x | R_a0x 5 | 91 | Num
Ra Table | R_a0x | R_a0x 6 | 106 | Num
Ra Table | R_a0x | R_a0x 7 | 112 | Num
Ra Table | R_a0x | R_a0x 8 | 107 | Num
Ra Table | R_a0x | R_a0x 9 | 121 | Num
Ra Table | R_a0x | R_a0x 10 | 147 | Num
Ra Table | R_a0x | R_a0x 11 | 186 | Num
Ra Table | R_a0x | R_a0x 12 | 269 | Num
Ra Table | R_a0x | R_a0x 13 | 748 | Num
Ra Table | R_a0x | R_a0x 14 | 7558 | Num
Calibration | Data | CC Gain | 10.445 | mohm
Calibration | Data | CC Delta | 10.424 | mohm
Calibration | Data | CC Offset | -1540 | num
Calibration | Data | Board Offset | 32 | num
Calibration | Data | Int Temp Offset | 0 | degC
Calibration | Data | Ext Temp Offset | 0 | degC
Calibration | Data | Voltage Divider | 7743 | mVolt
Calibration | Current | Deadband | 5 | mAmp
Security | Codes | Sealed to Unsealed | 36720414 | hex
Security | Codes | Unsealed to Full | ffffffff | hex
Security | Codes | Authen Key3 | 1234567 | hex
Security | Codes | Authen Key2 | 89abcdef | hex
Security | Codes | Authen Key1 | fedcba98 | hex
Security | Codes | Authen Key0 | 76543210 | hex