## Slide 1: Test Setup – Circuit Diagram

## Slide 2: Test Setup – Notes
Input Amplitude: 8 VPP
Input Frequency: 10 Hz, 40 Hz, 100 Hz
Input Source: Audio Precision, Unbalanced GND output
Samples Collected: 4,096
Expected Output: 8 VPP * [102.88 / (102.88 + 100k)] = 8.22 mVPP
Hardware: ADS1298RECGFE-PDK
Electrode Configuration:
  RA connected to P1 and IN2N
  LA, LL, V1 – V6 shorted to P2 and IN2P
  RLD connected to P6

## Slide 3: Test Setup – Register Settings

## Slide 4: S2 closed, 10 Hz, 8.23 mVPP

## Slide 5: S2 closed, 10 Hz, -39.75 dB

## Slide 6: S2 closed, 40 Hz, 8.16 mVPP

## Slide 7: S2 closed, 40 Hz, -39.83 dB

## Slide 8: S2 closed, 100 Hz, 7.83 mVPP

## Slide 9: S2 closed, 100 Hz, -40.19 dB

## Slide 10: S2 open, 10 Hz, 8.22 mVPP

## Slide 11: S2 open, 10 Hz, -39.76 dB

## Slide 12: S2 open, 40 Hz, 8.45 mVPP

## Slide 13: S2 open, 40 Hz, -39.52 dB

## Slide 14: S2 open, 100 Hz, 9.17 mVPP

## Slide 15: S2 open, 100 Hz, -38.81 dB