Experiment 1: Does the ADS7825 sample on the rising or falling edge of /BUSY?

Method:

- Vary the input voltage such that it has different values at the falling and rising edges of /BUSY

- Initiate conversion

- Compute the voltage equivalent of the ADC output code

- Compare voltage equivalent of code to the input voltage values at the rising and falling edges of /BUSY

Test conditions:

[A0, A1] = [0, 0]

SPI conditions:

/CS = “0” (shorted to GND)

R/nC width = ~2us

f_SCLK  = 5MHz

CPOL = 0, CPHA = 0   capture (MISO) on rising edge of SCLK

Delay first SCLK rising edge by 21us from falling edge of R/nC

Generate 18 SCLKs and capture MISO starting on 3rd rising edge of SCLK

Conversion #2

Conversion #1

-1.7V

+2.3V

300mV

300mV

300mV

Conversion #1 close-up:

Conversion #2 close-up:

Conclusion:

The results of Conversion #1 and Conversion #2 correspond to input voltage values that occurred at the FALLING EDGES of /BUSY. Therefore, the ADS7825 samples the input voltage at the falling edge of /BUSY.

Experiment #2: To which input channel does the output data belong?

Method:

- Apply different voltages to the input channels

- Set input channel address [A0, A1] to desired value

- Initiate conversion and compare voltage equivalent of ADC output code to input voltages

- Repeat steps (1) to (3) multiple times

Analog Input conditions:

AIN0 = 5V (shorted to power supply)

AIN1 = Hi-Z (1.79V when AIN1 selected, ~0V otherwise)

AIN2 = 2.5V (shorted to CAP pin)

SPI conditions:

/CS = “0” (shorted to GND)

R/nC width = ~2us

f_SCLK  = 5MHz

CPOL = 0, CPHA = 0   capture (MISO) on rising edge of SCLK

Delay first SCLK rising edge by 21us from falling edge of R/nC

Generate 18 SCLKs and capture MISO starting on 3rd rising edge of SCLK

Expected code out per channel:

Code(AIN0) = 16384 = 0x4000

Code(AIN1) = 5865 = 0x16E9

Code(AIN2) = 8192 = 0x2000

Timing diagram below shows typical conversion cycle; ADC powered up with [A0, A1] = [0, 0], AIN0 = 5V; code out = 0x4001

Latch conversion output data; Start acquiring (or “tracking”) next CH

Take Input sample; Latch next CH

Convert Input Sample

Delay to next convert signal >> 25us (very long sampling time)

MSB

LSB

Convert with no change to [A0, A1]  expected code = 0x4000, actual code = 0x4009; Next channel to convert = AIN0 (expect 0x4000 next)

Change [A0, A1] to [1, 0] and convert  expected code = 0x4000, actual code = 0x4005; Next channel to convert = AIN1 (expect 0x16E9 next)

Change [A0, A1] to [0, 1] and convert  expected code = 0x16E9, actual code = 0x16DA; Next channel to convert = AIN2 (expect 0x2000 next)

Change [A0, A1] to [0, 0] and convert  expected code = 0x2000, actual code = 0x2000; Next channel to convert = AIN0 (expect 0x4000 next)

Change [A0, A1] to [0, 1] and convert  expected code = 0x4000, actual code = 0x4006; Next channel to convert = AIN2 (expect 0x2000 next)

Change [A0, A1] to [1, 0] and convert  expected code = 0x2000, actual code = 0x1FFF; Next channel to convert = AIN1 (expect 0x16E9 next)

No change to [A0, A1] and convert  expected code = 0x16E9, actual code = 0x16DC; Next channel to convert = AIN1 (expect 0x16E9 next)

No change to [A0, A1] and convert  expected code = 0x16E9, actual code = 0x16DB; Next channel to convert = AIN1 (expect 0x16E9 next)

Conclusion:

The ADS7825 datasheet states that the output data captured in a given cycle corresponds to the input channel whose address was latched on the previous falling edge of /BUSY. We have verified that this statement is correct.