## Slide 1: ADC32RF45EVM Test with TSW14J10EVM and ZC706

## Slide 2
Test Setup:

## Slide 3
Hardware Setup:
Connect the three boards as shown in previous slide
Connect power to ZC706 and ADC32RF45EVM
Connect a USB cable between TSW14J10EVM and a host PC.
Connect a mini-USB cable between J1 of the ZC706 and a host computer that has Vivado installed.
Connect a 2GHz,12dBm clock source to J5 (ADC_CLK_IN) of the ADC32RF45EVM. Verify the board is configured for external clock mode.
Connect a 2GHz,12dBm clock source to J7 (LMK_CLKIN) of the ADC32RF45EVM. Verify the two clock sources are synchronized.
Provide a 600MHz, 6dBm tone to either AINM or BINP analog input SMA’s.

## Slide 4
Test Setup:
Single tone is given as input to the device. Data captured using HSDC Pro GUI
Test conditions:
Fs = external 2Gbps
LMK = external 2Gsps clock synchronized with Fs clock.
Fin = 600MHz
Mode 8224

## Slide 5: Open ADC32RFxx EVM GUI

## Slide 6: Click on “Low Level View” tab. Click on folder shown below to open configuration files to load.

## Slide 7: Navigate to “Bypass DDC Mode directory. Select “LMK_ADC32RF45_LMF_8224_ExtClock.cfg”

## Slide 8: Press board reset switch SW1. In Low Level View tab, navigate to Bypass DDC Mode directory and load file called ADC32RF4x_14bit_LMFS_8224.cfg”

## Slide 9: Click on the “LMK04828” tab then the “Clock Outputs” tab. Set CLKout 0 and 1 DCLK Divider to “4”. Unselect “Group Powerdown for CLKout 12 and 13, set the divider to 8 and set the DCLK Type to “LVDS”, as shown below.

## Slide 10
Move the file called “TSW14J10_ZC706_v2p8.bit”, located at:
C:\Program Files(86)\Texas Instruments\High Speed Data Converter
Pro\14J10ZC706 Details\Firmware to C:\ on the host computer running Vivado.

## Slide 11: Open Vivado. Click on “Open Hardware Manager”.

## Slide 12: Click on “Open Target”.

## Slide 13: Select “Open New Target”.

## Slide 14: Click on “Next”.

## Slide 15: Click on “Finish”.

## Slide 16: Click on “Program device the select the device that opens as shown below”.

## Slide 17: Click on the Bitstream file navigation box and select the file called “C:\TSW14J10_ZC706_2vp8.bit” Click on “Program”

## Slide 18: A new window will open showing the status of the programming. Once this reaches 100%, the board is programmed and ready to be used with the HSDC Pro GUI

## Slide 19: The three status LEDS labeled as “L C R” under the label “PL GPIO LEDS” should all be blinking now. They are located near the board power switch, as shown below.

## Slide 20
Open HSDC Pro GUI.
Click on “OK” to connect to the ZC706

## Slide 21
Click on the device drop down arrow and select “ADC32RF45_8224”.
Enter “2G” for the ADC Output Data rate.

## Slide 22
Click on “capture”.
The lane rate will be 10G and the FPGA reference clock will be 500M.

## Slide 23: 600MHz IF @ CH1

## Slide 24
600MHz IF @ CH2