## Slide 1: Software PACE detect with ADS1292
Tony Calabria
Texas Instruments

## Slide 2: Statement
PACE detections used to determine if a patient has a pacemaker.
PACE signal frequency is generally above 500Hz.
  Beyond the typical clinical ECG bandwidth of interest.
  Requires higher passband from digital filter  i.e. Increase data rate
Using a higher data rate increases bandwidth and, therefore can add to noise.
  Digital Filter used to help clean up ECG signal
All tests with the ADS1292ECG-FE Board and the Evaluation software available online.

## Slide 3: Bandwidth of interest
Including Gamma
EMG Signal
150Hz
500Hz
60Hz
20Hz
ECG Signal
120Hz
EEG Signal
10 KHz
Pacemaker Signal
Respiration 0 to 0.4Hz
Frequency
Max Signal Amplitude
~5mV
~5mV
~1mV
~10mV
*Amplitude dependent on external current source
*Frequency ranges are estimates. Actual frequency range will depend on specific application
50Hz
0.05Hz

## Slide 4: ECG Signal at 500SPS (no filtering)

## Slide 5: ECG Signal at 8kSPS (no filtering)

## Slide 6: ECG Signal at 8kSPS with Digital Filter
100Hz Digital LPF enabled

## Slide 7: Side by Side Comparison
500SPS no digital Filter
8kSPS 100Hz digital Filter
Comparable ECG pulses using 500SPS and 8kSPS (w digital filter)
Running at 8kSPS will be sufficient bandwidth to monitor PACE

## Slide 8: Enabling PACE
All tests with the ADS1292ECG-FE Board and the Evaluation software available online.
Data Rate set to 8kSPS for PACE detection
  Digital Filter cutoff frequency is 2kSPS.
  10,000 samples used
Ventrical Pacer option enabled in Fluke Medsim 300B for test
  Control over Ventrical Pace signal width and amplitude
    Amplitude Range: -100mV to 100mV
    Signal Width Range: 0.1ms to 2 sec

## Slide 9: Ventrical PACE Test 1
Simulator set to 2mV Amplitude and 1ms pace signal width

## Slide 10: Ventrical PACE Test 1 Zoomed

## Slide 11: Ventrical PACE Test 2
Simulator set to 10mV Amplitude and 0.1ms pace signal width

## Slide 12: Ventrical PACE Test 2 Zoomed