## User Guide
Yellow marked cells require user input
Green marked cells are automatically calculated and should not be changed manually

## Convert Code to Voltage
Conversion Calculator
Codes to Voltage
ADC Code | 1 | Hex
| 1 | Decimal
VREF | 2.048 | V
PGA Gain | 1 | V/V
1 LSB | 244.1406 | nV
Diff Input Range | 2.048 | +/- V
Vout | 2.4414e-07 | V
Conversion Calculator
Voltage to Codes
Voltage | 1 | V
VREF | 2.048 | V
PGA Gain | 1 | V/V
1 LSB | 244.1406 | nV
Diff Input Range | 2.048 | +/- V
ADC Code | 4096000 | Decimal
| 3E8000 | Hex

## PGA - Common-Mode Voltage
| Value | Unit
AVDD | 3.3 | V
AVSS | 0 | V
AINp | 1.28 | V
AINn | 1.25 | V
PGA Gain | 32 | V/V
VREFp | 1.25 | V
VREFn | 0 | V
VIN = (AINp - AINn) | 0.03 | V | Cells will turn red when VCM or VIN are not within allowed limits.
VCM = (AINp + AINn)/2 | 1.265 | V
VCM Lower Limit | 0.58 | V
VCM Upper Limit | 2.72 | V

## 4-Wire RTD
ADS1248 PARAMETERS
Resolution | 24 | Bits
AVSS | 0 | V
AVDD | 3.3 | V
Sampling Rate | 5 | SPS
PGA GAIN | 32 | V/V
Input Referred Noise | 0.32 | µVpp | Look up according noise values in the ADS1248 datasheet for the repsective gain and data rate setting
RREF | 5 | kΩ
IDAC | 250 | µA
RTD CALCULATIONS
Nominal Resistance | 100 | Ω
Sensitivity | 0.385 | Ω/°C
| Temperature [ᵒC] |  | Resistance [Ω] | VRTD [V]
RTD RANGE | Typ | 50 | 119.25 | 0.029812
| Min | 0 | 100 | 0.025
| Max | 100 | 138.5 | 0.034625
Resolution
VREF | 1.25 | V
PGA x VRTD(max) | 1.108 | V | Tune PGA Gain, IDAC value and RREF so that this value is close to, but still below VREF
LSB Size | 0.0046566 | µV
Resolution of 1 Code | 4.838e-05 | °C
Noise Free Resolution | 0.0033247 | °C
IDAC Compliance Check
IDAC Compliance Voltage | 2.4 | V
V_AINp (max) | 1.2846 | V | If this cell turns red your IDAC will be out of compliance and you need to reduce the IDAC or RREF value
Common-Mode Voltage Check
VCM = (AINp + AINn)/2 | 1.2673 | V | If this cell turns red your setup does not meet the common-mode voltage requirement
VCM Lower Limit | 0.654 | V
VCM Upper Limit | 2.646 | V